Nano-inXider SOP

Important information

Microfocued source 40um

Copper anode 8KeV

Beam shape – 700um full beam

Sample to detector distance 938mm (standard – can be adjusted) based on reference plane of 0.5mm above standard sample holder

Detector angle 0*

Vacuum 0.1mbar

X-Z translation 35mm

Psi rotation 190* (allows for higher azimuthal ranges for virtual detector)

2 x 12.5um Kapton windows (one is removable)

Tensile stage available  (0-200N) with 0.01N resolution (haven’t got this as humidity is a problem in vacuum)

Temperature stage -150*-250*

|  | Beam size | flux | qmin |
| --- | --- | --- | --- |
| VHF | 1mm | 100Mph/s | 0.011 |
| VHR | 200um (pinhole) | 4Mph/s | 0.0029 |

Secret0 for Superuser

User / Password = Nano / Xenocs

Always shut WAXS detector when opening

Scan in sample alignment Range 5mm (-2.5 – 2.5) either side of point

60s for calibrants

LaB6 (for WAXS detector)

AgBe (for SAXS detector)

Glassy Carbon for Absolute intensity

Empty hole for beam alignment

- Log into computer (User Nano, Pass Xenocs)

- Open both the NanoinXider software and XSact software (user Himadri)

- Clicking ‘start’ will bring it to vacuum

- Can change the stage by clicking change sample stage

  - Make sure the WAXS detector is shut

  - Page 44 and 69 of manual

- Set folder for data output

- Put in sample

- Pump to vacuum (it moght fail but just run it again)

- Scan intensity range (5mm and 26 pts no pts plus 1)

- ‘0’ file SAXS, ‘1’ file WAXS